# D-Control

# Contents

Contents	1

1	Installation	2

2	Application Description 	2

2.1	Adding AC source	3

2.2	AC Source Control Panel	3

2.3	Voltage Switching Program	4

2.4	Voltage Pattern Diagram	4

2.5	Communication 	5

2.6	About the Program	5

3	Details of AC 250Kxxx Source Communication	6

3.1	AC250Kxxxx Source Commands	6

3.2	BCC Checksum Generation:	7

# Installation

The„D-Control“ application comprises an independently startable file „dcontrol.exe“. We recommend creating a new folder, C:\Program Files\DControl, and instal the D-control software into it. At the first start the „dcontrol.ini“ file will be created in the same directory as the software itself. This is also the place where measured data will be stored unless a different directory is selected.

# Application description

After a start-up and the initial warning confirmation the application form pops up. On the left side you will see global buttons described in the picture.

Close application

Change language

Connect

AC source

Turn on outputs on all AC sources

Turn off outputs on all AC sources

Set all connected AC sources to zero voltage

In the form you will find four bookmarks operable by function keys or your computer mouse: F1 – Overview of the attached AC sources, F2 – Voltage diagram, F3 – Communication between the application and AC sources, F8 – About the program.

## Adding AC source

Press „Connect“ for adding an AC source and launching communication with it. Proceed by choosing a communication line and entering a serial port through which the AC source will be connected to your computer. Enter the address of the source and the colour of the source’s control panel. After this press „Add“ to create a control panel for the AC source in the overview.

If you wish to add another AC source through a serial line, choose another free communication line and set a free serial port in your computer. AC sources using the RS485 communication interface can be connected by means of a RS485/RS232 per each communication line. In such case, the AC sources must have different addresses.

## AC source control panel

The control panel enables you to set the desired voltage, adjust the stability tolerance [1 .. 9 V], as well as turn the output on and off. Also you can turn the stabilizer off, as well as set and check the program for automatic voltage change.

AC source type and serial No.

Voltage setting

Current voltage

Allowed tolerace value

Start / Stop the program and show its state

Record voltage pattern into a file

Stabilisation tolerance setting

Setting the program for automatic voltage change

Changing the panel colour

Removing the panel and terminating the communication with the AC source

Turn voltage output on / off

Turn voltage stabilisation on / off

The voltage (or tolerance) is set by entering the desired value into the respective space and pressing Enter.

To tun the device on/off click repeatedly on „Out“. Similarly, for turning the stabilization on/off click on „Sta“.

## Voltage Switching Program

The Voltage Switching Program consists of several steps that correspond with the lines of the table comprising the following values:

- time [s] specifies the waiting time in the step

- U [V] voltage set in the step

- Setting the next step, i.e. the row to which the program will proceed

A new row in the table is created by entering the “Next” value. Upon pressing the “Enter” key in the “Next” field, the cursor moves to the “Time” field of a particular step. In case you enter value 0 to the “Next” step field, the application stops. If you enter value 1 to the “Next” field, or any other value lower than that of the current row, the application will cycle in an infinite loop until being interrupted by the user. This is convenient to perform long-term tests of devices connected to the source.

The “Save” button is used to store the program for reuse in the future. The “Load” button is used to open a program stored in a file. The program editing window is closed by clicking the “OK” button.

## Voltage Pattern Diagram

The “F2 – Diagram” tab is used to view graphical voltage pattern of a particular source.

The list of all the attached AC sources. The ticked-off AC sources are marked in the diagram

Print the diagram

Number of values on the „x“ axe

The „y“ axe range

Parametres of the diagram display:

3D / 2D choice

3D perspective

enlargement

rotation

## Communication

The “F3 – Communication” tab is used to view the course of communication at individual lines.

## About the program

The last tab shows information about the D-Control application along with the license agreement. At the first launch of the application you are required to read the complete license agreement and press the “I Accept” button to confirm the provisions of this agreement.

# Details of AC 250Kxxx Source Communication

Data Communication Setup:

9600Bd. 8bit, No parity, 1 start/stop bit

Packet pattern:

@ A A P P P X X X …. CR

Each packet starts with “@” character. Then, a two-byte address follows (AA). An address is shown in HEX in the text mode. The range for addresses is from 00 to 1F, i.e. total of 32 addresses. A packet with “FF” address is interpreted as a command for all devices, while devices do not reply to these packets. Then there is a three-byte command. It must not contain @ and # characters. Command is followed by parameters. Parameters must not contain @ and # characters. Each packet is terminated by a BCC checksum (two bytes) and CR character ($0D).

Devices respond immediately upon receiving a complete packet with a valid address.

Response examples:

# A A P P P X X X …. CR

or

# A A O K CR

or

# A A Err CR

or

# A A T Y P X X X X X X X X X X V C X X X X CR

## AC250Kxxx Source Commands

1) 	inquiry about source type and its serial No.	@0AID?3D$0D

2) 	change voltage to 255		@0ANAP255EC$0D

3) 	change voltage to 000		@0ANAP000E0$0D

4) 	change voltage to 100		@0ANAP100E1$0D

5) 	check the voltage value	@0ANAP???0D$0D

6) 	turn on the output		@0AOUT19A$0D

7) 	turn off the output		@0AOUT099$0D

8) 	inquiry about output state	@0AOUT?A8$0D

9) 	stab. OFF			@0ASTA089$0D

10)	stab. ON			@0ASTA18A$0D

11)	inquiry about stab. state	@0ASTA?98$0D

12)	inquiry about tolerance	@0ATOL?9F$0D

## BCC Checksum Generation:

1) Convert source address to HEX format (source address 10 is converted to packet address 0A).

2) Make a dec sum of individual bits converted to ASCII, located before the checksum (without the @ character!).

3) Keep subtracting 256 from the result, until the final value is less than or equal to 256

4) Convert the result to HEX.

5) Now you have obtained the checksum (BCC).

Example:

@01ID?2D$0D

PAKET	HEX

0		48

1		49

I		73

D		68

?		63

_____________

Sum		301

-256

_____________

Result: 	45 => 2D

Example:	@01NAP000D0$0D		To send a command to change voltage to 0V

@0AID?3D$0D	To send an inquiry about source type and its serial No.